Development and Characterization of a Paper-Based Graphite Sensor Using NaCl Electrolyte

1. Sensor Fabrication

Pencil Graphite-Based Sensor

- Substrate: Grade 1 filter paper (hydrophilic)

- Active area:

  - Width = 0.4 cm

  - Height = 2.54 cm

- Electrode gap: 0.1 cm

- Conductive material: Graphite (manually deposited)

Electrical properties

The resistance of the graphite sensor follows:

R = ρLA

where,

- R = resistance

- ρ = resistivity of graphite network

- L = effective length of the graphite track

- A = effective cross-sectional area

Observations:

- Resistance is directly proportional to sensor length

- Resistance is inversely proportional to the amount (weight) of graphite deposited

- Initial dry resistance range: 6 kΩ – 22.3 kΩ

The paper substrate absorbs the electrolyte, enabling ionic–electronic mixed conduction through the graphite network.

Graphite INK-Based Sensor

- Substrate: Grade 1 filter paper (hydrophilic)

- Active area:

  - Width = 0.4 cm

  - Height = 2.54 cm

- Electrode gap: 0.1 cm

- Conductive material: Graphite INK (manually deposited)

Electrical properties

The resistance of the graphite sensor follows:

R = ρLA

where,

- R = resistance

- ρ = resistivity of graphite network

- L = effective length of the graphite track

- A = effective cross-sectional area

Observations:

- Resistance is directly proportional to sensor length

- Resistance is inversely proportional to the amount (weight) of graphite deposited

- Initial dry resistance range: 50Ω – 1.7 kΩ

The paper substrate absorbs the electrolyte, enabling ionic–electronic mixed conduction through the graphite network.

Pencil Graphite-Based single Probe Sensor

- Substrate: Grade 1 filter paper (hydrophilic)

- Active area:

  - Width = 1 cm

  - Height = 2.54 cm

- Conductive material: Pencil Graphite (manually deposited)

Electrical properties

The resistance of the graphite sensor follows:

R = ρLA

where,

- R = resistance

- ρ = resistivity of graphite network

- L = effective length of the graphite track

- A = effective cross-sectional area

Observations:

- Resistance is directly proportional to sensor length

- Resistance is inversely proportional to the amount (weight) of graphite deposited

- Initial dry resistance range: 2.6 kΩ – 14.8 kΩ

The paper substrate absorbs the electrolyte, enabling ionic–electronic mixed conduction through the graphite network.

Copper-Based Sensor (Reference Probe)

Electrode material: Copper plates

Dimensions: Identical to graphite sensor (0.5 × 2.54 cm)

- Substrate: Hydrophobic (does not absorb water)

- Initial resistance: Negligible (≈ 0 Ω)

This probe serves as a reference conductivity sensor, where conduction is dominated purely by the electrolyte.

2. Experimental Setup

Electrical Configuration

- Excitation voltage: 5 V DC

- Measurement circuit: Voltage divider

Vout= Vin×RfixedRsensor+Rfixed

Where:

- Rfixed= 10kΩ

- Rsensor varies with NaCl concentration

The output voltage is proportional to the effective conductivity of the sensor–electrolyte system.

Copper Sensor Setup					Graphite Sensor Setup

Data Acquisition Hardware

- ADC: ADS1115 (16-bit)

- Gain setting:

- ads.setGain(GAIN_TWOTHIRDS);

  - Input range: ±6.144 V

  - Resolution: 1LSB = 0.1875mV

- Controller: Arduino UNO (ATmega328)

ADC counts are converted to voltage as:

V = ADCcount×0.1875 mV

3. Data Collection Protocol

- Base solvent: 20 ml distilled water

- Solute: 99.9% pure NaCl

- Concentration range:

  - DRY_RUN (no solution)

  - 0 gm to 11 gm NaCl in 20 ml water.

- Procedure:

  - Sensor exposed to solution.

  - Continuous ADC data recorded.

  - Same protocol used for graphite and copper sensors.

Raw Data Plotting

Data From Graphite sensor

Data From Copper Sensor

4. Data Screening and Processing

Noise Reduction: Moving Average

A moving average filter is applied:

ysmoothn= 1WK=-W2W2yn+k

- Window size: 1000 samples

- Purpose:

  - Suppress high-frequency noise

  - Remove transient spikes

  - Extract steady-state response

4.2 Feature Extraction (Per Concentration)

From the smoothed signal, the following parameters are computed:

- Minimum value:

ymin=minysmooth

- Maximum value:

ymax=maxysmooth

- Mean value:

μ=1N∑ysmooth

- Standard deviation:

σ= 1N∑ysmooth- μ2

- Range:

Range = ymax-ymin

his reduces large time-series data into physically meaningful descriptors.

Min–Max Envelope Analysis

For each concentration:

- Min and max values form a response envelope

The average envelope is defined as:

yavg=ymin+ymax2

This represents the effective sensor response while preserving variability.

Linearity Evaluation

Linear regression:
y = mx + c

Coefficient of determination:

R2= 1 -∑y - yfit2∑y - y2

Computed separately for:

- Min curve

- Max curve

- Average curve

This quantitatively evaluates concentration–ADC linearity.

5. Graphs After Screening

Pencil Graphite Based 2 probe sensor

Graphine INK Based 2 probe sensor

Pencil Graphite Based single probe sensor

Copper based 2 probe sensor

6. Key Findings

- Monotonic response:
ADC values increase consistently with NaCl concentration for both sensors.

- Copper sensor:

  - Higher linearity

  - Narrow response envelope

  - Stable and repeatable behavior

- Graphite-paper sensor:

  - Wider envelope due to material and substrate effects

  - Clear linear trend in averaged response

  - Functional as a low-cost sensing element

- Data processing pipeline successfully:

  - Removes noise

  - Preserves physical trends

  - Enables quantitative comparison

GitHub Link:

https://github.com/RaunakR16/Paper_Sensor/tree/main